Absorption Spectroscopy
BEER LAMBERT LAW:
dI over dl ~=~ I C %epsilon'

---
BEER LAMBERT LAW:
dI over dl ~=~ I C %epsilon'
ln left (I_0 over I right ) ~=~ C %epsilon 'l
A( %lambda) ~=~ log_10 left (I_0 over I right ) ~=~ C %epsilon (%lambda)l
A is the absorbance or optical density (OD) of the sample at a particular wavelength.
This measure can be used to determine the concentration of sample
The extinction coefficient ε is a function of wavelength and independent of concentration, where ε = ε'/2.303. ε is typically expressed in OD/mol for a 1 cm pathlength.

---
Double-beam spectrophotometer suitable for measuring difference spectra. This can be useful for simply subtracting the background absorbance of the buffer or the cuvette in order to correct for the extra absorbance, but it cal also be used for detecting small changes in absorbance in difference absorption experiments, for example when enzyme and substrate form a complex. Their combined absorption is measured in series in the top arrangement of two separate cells, while the lower beam measures the absorption of the complex. If the concentration and pathlength is identical, then the only difference would be due to absorbance changes when the complex is formed.

---
Why is this plot not linear for the entire range?
At some point the absorbance is so high that not sufficient light passes through to the detector, and linearity is no longer satisfied.
The absorbance at which an instrument becomes non-linear depends on the following factors:
Concentration of the analyte
Lamp intensity at the measured wavelength
Extinction coefficient of the analyte at measured wavelength
Sensitivity of detector at the measured wavelength
Relationship between
Absorbance and Concentration
For best accuracy, always measure between 0.1 – 1.0 OD

---
Useful Tips for Absorbance Spectroscopy
Modern spectrophotometers typically store a reference spectrum in memory and perform the subtraction internally to report true analyte absorbance.
Background absorbance results from cuvette and buffer.
For UV mneasurements, only use quartz cuvettes. Glass and plastic cuvettes absorb in the UV and can only be used in the visible range.
Many buffer components absorb in the UV, in particular beta mercapto-ethanol and di-thiothreotol. TCEP can be used down to 240 nm.
Many buffers absorb to varying degrees below 240 nm, including HEPES, TRIS, acetate, MOPS; sodium and potassium phosphate buffers do not absorb and can be used to quite low wavelengths.
Always consider the total absorbance when subtracting buffer absorbance to make sure you are not measuring outside the dynamic (linear) range of the spectrophotometer.

---
Keep in mind the logarithmic nature of absorbance:
1.0 OD means that only 10% of the light is seen by the detector, 90% is absorbed. At 2.0 OD only 1% of the light reaches the detector, 99% is absorbed.
Some problems:
What is the absorbance when 50% of the light is absorbed by the sample?
What is the molar concentration of protein that absorbs 0.75 OD at a wavelength where its extinction coefficient is 24,300 OD/(mol*cm)? Assume a 1 cm pathlength cell.
A ~=~ - log_10 left ( I over I_0 right) ~=~ - log_10 left ( 0.5 right) ~=~ 0.301 OD
A ~=~ C %epsilon l ~~~C ~=~ A over {%epsilon l} ~=~ 0.75 over 24,300 ~=~ 30.86 %mu M

---
Absorption Spectroscopy of Biopolymers:
For biologists, the absorption properties of nucleic acids and proteins are of most interest. Nucleic acids and proteins have distinct chromophores that give rise to unique absorption curves. Typically, proteins and nucleic acids have chromophores that absorb in the UV (< 300 nm). Some proteins have chromophores in the visible (e.g. GFP, heme proteins, or proteins complexed with transition metals that give rise to unique absorbance spectra). Measurements are typically performed in aqueous solvents, which limits measurements to wavelengths > 170 mn. Buffers, salts etc. need to be chosen with care to avoid background absorbance.
For proteins, the following chromophores are dominant:
Peptide bond: delocalized π electrons (carbonyl-nitrogen) n to π* transition (weak intensity in the far-UV at 210-220 nm)
Peptide bond: delocalized π electrons (carbonyl-nitrogen) π to π* transition (very strong intensity in the far-UV at ~ 190 nm)
Amino acids with aromatic side chains that have delocalized π electrons, in order of absorbance strength (λmax weakest to strongest): Phe (258 nm), Tyr (274 nm), Trp (280 nm).
Disulfide linkages from oxidized sulfur atoms on Cys residues (250-270 nm), as well as His (both are weak).

---
For proteins without prosthetic groups, the number of tryptophan and tyrosine residues can be used to estimate protein concentration by measuring the absorbance of the protein at 280 nm. The molar extinction coefficients of Trp and Tyr are 5,700 and 1,300 OD 280 nm/(mol*cm), so the molar extinction coefficient of the protein can be estimated based on the number of tryptophan and tyrosine residues (nTrp and nTyr) with this formula:
To convert into mg/ml, divide by the molar mass, M:
A more sensitive method, due to the typically higher extinction, is to measure the protein concentration by monitoring the absorption at 230 nm. Here the extinction is roughly 300 OD 230/(mol*cm) for each peptide bond.
%epsilon_{protein, 280 nm} ~=~ 5,700 nTrp `+` 1,300 nTyr
C_{protein, 280 nm} ~=~ {5,700 nTrp `+` 1,300 nTyr} over M

---
For nucleic acids, the strong extinction of the purine and pyrimidine bases contributes to their absorption in the near UV at around 260 nm, and in the far UV between 190-210 nm.

---
Like phosphate buffer, the phosphate groups do not contribute to the extinction.
Within experimental error, ATP, ADP and AMP all have the same extinction properties. Absorbance of an individual nucleoside at 260 nm is about 10,000 OD 260 nm/(mol*cm). This strong extinction allows measurement down to concentrations of ~ 3 ug/ml.
Contributions to the absorbance spectrum by the environment:
Dipole interactions with the solvent, if the solvent is polar like water, can be significant, and can change the absorption characteristics of chromophores when the solvent is changed to a non-polar solvent. Similarly, if the environment of chromophore changes within the molecule (e.g., a surface solvent-exposed tryptophan becoming buried in a hydrophobic interior of the protein), extinction intensities and absorption bands change. This can be used to monitor changes of a macromolecule, such as in:
melting studies
pH changes
oxidation states
Conformational changes such as unfolding
Solvent differences
Complex formation

---
Hypochromism is the loss of absorption intensity in some regions of the absorbance spectrum.
Hypochromism is caused by the alignment of transition dipoles.
Examples for the effect of hypochromism can be observed most often during binding or aggregation. A very good example is the 30% increased absorbance DNA nucleosides over single stranded to double stranded DNA. The alignment of the bases (stacking) can reduce the absorbance.
Hypochromism depends on 1/r3 of the aligned residues giving rise to the dipoles

---
Deconvolution of multiple basis spectra to quantify individual components:
When known absorbance spectra are sufficiently different and hypochromism is negligible, the individual components can be mathematically separated to determine relative amounts of each constituent.
F_{ base }~=~ SUM G_i left ( %sigma_i, %lambda_i, c_i right ) ~=~ SUM c_i e ^{{-(x-%lambda_i)^2}over{2 %sigma_i^2} }
F_mix ~=~ sum from{i}a_i F_{base,i}

---
Experiment: Prepare food color solutions from primary colors, and measure their pure absorbance spectra. Take multiple concentrations and globally fit each color to a sum of Gaussian terms. Then mix the colors with known ratios and try to determine the composition by performing a linear decomposition of the summation spectra into their basis spectra.